Ventilation Patterns

| Ventilation-Pattern | Settings | Mandatory/ Optional | Possible Units | Notes |
| --- | --- | --- | --- | --- |
| CMV | inflation-type | M |  | only one inflation-type |
|  | rate | M | min-1 |  |
|  | BAP | M | cm H2O, kPa |  |
|  | low pressure limit | M | cm H2O, kPa |  |
|  |  | O |  |  |
| A/C | inflation-type | M |  | only one inflation-type |
|  | rate | M | min-1 |  |
|  | BAP | M | cm H2O, kPa |  |
|  | flow trigger level | O | l∙s-1, l∙min-1 |  |
|  | pressure trigger level | O | cm H2O, kPa |  |
|  | neural trigger level | O | mV |  |
|  | inflation trigger algorithm | O |  |  |
|  | low rate limit | M | min-1 |  |
|  | high rate limit | M | min-1 |  |
|  | low minute volume limit | M | l∙min-1 |  |
|  | high minute volume limit | O | l∙min-1 |  |
| SIMV | inflation-type(s) | M-assured O-support inflation-type |  | one assured inflation-type ± one support inflation-type |
|  | rate | M | min-1 |  |
|  | synchronization window time | M | ms, s |  |
|  | flow trigger level | O | l∙s-1, l∙min-1 |  |
|  | pressure trigger level | O | cm H2O, kPa |  |
|  | neural trigger level | O | mV |  |
|  | BAP | M | cm H2O, kPa |  |
|  | low rate limit | M | min-1 |  |
|  | high spontaneous breath rate limit | M | min-1 |  |
|  | minimum minute volume limit | O |  |  |
| S/T ventilation | support inflation-type | M |  |  |
|  | flow trigger level | O |  |  |
|  | pressure trigger level | O |  |  |
|  | neural trigger level | O |  |  |
|  | assured inflation-type | M |  |  |
|  | timed-ventilation rate | M |  |  |
|  | BAP | M |  |  |
|  | low rate limit | M |  |  |
|  | high spontaneous breath rate limit | M |  |  |
|  | low minute volume limit | O |  |  |
|  | high minute volume limit | O |  |  |
| CSV | inflation-type | O |  |  |
|  | flow trigger level | O |  |  |
|  | pressure trigger level | O |  |  |
|  | neural trigger level | O |  |  |
|  | BAP | M |  |  |
|  | high spontaneous breath rate limit | M |  |  |
|  | low rate limit | M |  |  |
|  | high rate limit | M |  |  |
|  | apnea ventilation | M or O if backup ventilation enabled |  |  |
|  | apnea ventilation inflation-type and apnea ventilation rate | M if enabled |  |  |
|  | backup ventilation | M or O if apnea ventilation enabled |  |  |
|  | backup ventilation inflation-type and backup ventilation rate | M if enabled |  |  |
| CPAP | BAP | M |  | (inflation-types not permitted) |
|  | low rate limit | O |  |  |
|  | low minute volume limit | O |  |  |
|  | apnea ventilation | M or O if backup ventilation enabled |  |  |
|  | high spontaneous breath rate limit | M |  |  |
|  | apnea ventilation inflation-type and apnea ventilation rate | M if enabled |  |  |
|  | backup ventilation | M or O if apnea ventilation enabled |  |  |
|  | backup ventilation inflation-type and backup ventilation rate | M if enabled |  |  |
| Bilevel AV | Phigh or BAPH | M |  |  |
|  | Plow or BAP | M |  |  |
|  | inspiratory time or Thigh | M |  |  |
|  | rate or Tlow | M |  |  |
|  | I:E ratio | O |  |  |
|  | high spontaneous breath rate limit | M |  |  |
|  | high minute volume limit | M |  |  |
|  | low minute volume limit | O |  |  |
|  | concurrent support inflation-type | O |  |  |
| APRV | Phigh or BAPH | M |  |  |
|  | Plow or BAP | M |  |  |
|  | inspiratory time or Thigh | M |  |  |
|  | rate or expiratory time or Tlow | M |  |  |
|  | I:E ratio | O |  |  |
| Bilevel PAP | Phigh or BAPH | M |  | NIV |
|  | Plow or BAP | M |  |  |
|  | inspiratory time or Thigh | M |  |  |
|  | rate or Tlow | M |  |  |
|  | I:E ratio | O |  |  |
|  | high spontaneous breath rate limit | M |  |  |
|  | high minute volume limit | M |  |  |
|  | low minute volume limit | O |  |  |

Mode=ventilation pattern + one or two inflation-types depending on ventilation pattern + FiO2  ± TC ± ACAP xor ( ± ACAPH xor ACAPL )  ± breathing system compliance compensation  ± breathing system leakage compensation  ± leak compensation

Where xor is one or the other not both

| Inflation Type | Attribute Settings | Mandatory/Optional | UOM |
| --- | --- | --- | --- |
| volume control VC | tidal volume | M | ml |
|  | inspiratory flow rate | M | l/m |
|  | rise time | O | ms, s, % of inspiratory time |
|  | flow pattern | O | square wave, ascending ramp, descending ramp |
|  | inspiratory time | M or O if I:E ratio set | s |
|  | inspiratory-pause time | O | ms, s |
|  | I:E ratio | O |  |
|  | high airway-pressure limit or high pressure relief limit | M | cm H2O, kPa |
| pressure control PC | inspiratory pressure | M | cm H2O, kPa |
|  | inspiratory time | M or O if I:E ratio set | ms, s |
|  | I:E ratio | O |  |
|  | rise time | M | ms, s, % of inspiratory time |
|  | high tidal volume limit | M | ml |
|  | high airway-pressure limit or high pressure relief limit |  |  |
| volume targeted pressure control vtPC | tidal volume (target) | M | ml |
|  | rise time | M | ms, s, % of inspiratory time |
|  | inspiratory time | M or O if I:E ratio set | ms, s |
|  | I:E ratio | O |  |
|  | high airway-pressure limit or high pressure relief limit | M | cm H2O, kPa |
|  | low tidal volume limit | O | ml |
| pressure support PS | inspiratory pressure or Δ support pressure Δ inspiratory pressure | M | cm H2O, kPa |
|  | inspiratory rise time | O | ms, s |
|  | inspiratory-termination flow | M | ms, s, % of inspiratory peak flow |
|  | high inspiratory time limit | M | ms, s |
|  | low tidal volume limit | O | ml |
|  | high airway-pressure limit or high pressure relief limit | M |  |
| volume support VS | tidal volume (targeted) | M | ml |
|  | inspiratory rise time | O | ms, s, % of inspiratory time |
|  | inspiratory-termination flow | M | % of inspiratory peak flow |
|  | high airway-pressure limit or high pressure relief limit | M | cm H2O, kPa |
|  | high inspiratory time limit | M | ms, s |
|  | low tidal volume limit | M | ml |
| effort support ES | effort support % | M | % |
|  | high tidal volume limit | M | ml |
|  | high airway-pressure limit or high pressure relief limit | M | cm H2O, kPa |
|  | maximum inspiratory flow | M | l/m |