# Routes and Lines

A route of administration refers to the route by which a drug or other substance enters the body.  It doesn’t define any physical device that may be used although the general type of device can often be inferred from the route e.g. subarachnoid implies the use of a spinal needle of some type and possibly a catheter.

Some routes have implicit anatomical references e.g. nasal, rectal so there’s no need to supply these separately. However, a majority of descriptions of routes  don’t include any information about the specific anatomical site of administration e.g. ‘Intravenous’ just says ‘the drug goes into a vein’ not which vein.

A line is a real physical system by which a drug is administered. It follows that lines always have specific, identifiable anatomical sites e.g. [left] cephalic vein (one of the veins in the forearm), [right] Proximal tibia (leg bone used for intra-osseous drug administration). So if a line is placed e.g. an IV line or epidural then the specific anatomical details are known and from these the route can be inferred. However, if we start with a route the anatomical site has to be added, it cannot be inferred, except in a few cases where the route also defines the anatomy as mentioned above.

There are many classifications of routes that share important features but then differ in detail. For example, if a drug is given sub-lingually (placed under the tongue and allowed to dissolve) then, technically, the route is transmucosal  because the drug is absorbed through the mucous membrane, it just happens that the site of that membrane is the mouth. So the ‘correct’ way to classify the route would be to say firstly that is is transmucosal and then to define the site as sub-lingual. However, some authors classify it as an oral route just because the drug is placed in the mouth. This is perhaps intuitive but it’s wrong because the oral route means that the drug is taken via the mouth i,e, it’s placed in the mouth and then swallowed.  The classification of drug routes given in Wikipedia is different to that used by the FDA and both differ from SNOMED CT and MESH.

Two scenarios for drug administration are envisaged:

- A drug is given via a line of some kind (all lines are by definition a  Parenteral route)

- A drug is given by some other means e.g. tablet, injection, suppository, patch etc.

Intra-operatively most drugs will be given through a line or by injection

Thus there should be options for:

- Route

- Associated line

Case 1: A drug is given via a line of some kind

In the drug / fluids dialogues there should be a means to associate a drug or fluid with a line and in the case of multi-lumen lines, with the specific lumen of that line. This means that there is no need separately to record details of the route or anatomical site as these are already linked to the chosen line.

Case 2: A drug is given by some other means e.g. tablet, injection, suppository, patch etc.

There should be a means to record the route of administration that will include:

- The type of route e.g. Intramuscular (I.M.) injection

- The anatomical site of administration where relevant e.g.[left] deltoid

Note that the general nature of the route, could be classified by an application.

# Lines

Not all types of line are used for drug administration. Some lines are used to measure, for example, blood pressure and to allow samples to be taken (intra-arterial lines can be used to sample blood for blood gas analysis). Drugs are most often administered using Peripheral Venous Catheters (PVCs), some drugs may be administered using a Central Venous Catheter (CVC)  but drugs are very rarely administered using intra-arterial lines as there are significant risk factors when this is done.

## Multi-lumen lines

Some lines have more than one lumen i.e. the line comprises more than one tube so that each tube can be used independently. In this case the number of lumens needs to be known and when a drug administration or any other procedure is associated with the line, there should also be a means to identify the specific lumen.  Multi-lumen lines are generally central lines. Each lumen has an associated port at it’s ‘open’ end i.e. the end away from the patient.

Peripheral lines are not usually multi-lumen. However, it is common to have a connector that allows a drip to be attached and a separate injection port so bolus drugs can be given.

## Lines recording options

Type of line:

- Peripheral Venous Catheter (PVC) or (PIVC)

- Central Venous Catheter (CVC)

- Peripherally Inserted Central Venous Catheter (PICC)

- Arterial (A-line)

- Intra-Osseous

- Implanted Portal Catheter

## Generic recording requirements for all lines

- Skin Preparation

- Aseptic technique

- Fixation (method by which the line is secured and sealed)

- Problems / issues

More information on specific recording requirements for each type of line are shown below:

## Central Venous Catheter (CVC) and Peripherally Inserted Central Venous Catheter (PICC)

### Common Insertion Sites for CVC

- Internal Jugular

- Subclavian vein

- Axillary vein

- Femoral vein

- Other [specify]

### Common Insertion Sites for PICC

- Cephalic vein

- Basilic vein

- Brachial vein

- Other [specify]

### Catheter Attributes

- Non-tunneled (e.g. Quinton)

- Tunneled (e.g. Hickman, Groshong) – strictly speaking the tunneling is a description of method not an attribute of the line

- Gauge (usually expressed as French gauge)

- Length (cm.)

- Number of lumens (1 – 5)

- Cuffed

### Adjunctive technique

- Ultrasound

- Chest X-ray (CXR)

- Flouroscopy

- Other [specify]

### Problems / issues

- Inadvertent arterial puncture

- Venous tear

- Haemothorax

- Pneumothorax

- Other [specify]

Note: tunnelled catheters are actually two separate components. One is the actual venous catheter and is introduced in the normal way. The placement of the tunnelled portion is done as follows:

- An exit site is identified about 10 cm. from the insertion point of the venous catheter and a hole is made with a scalpel

- A tool (e.g. pituitary clamp) is pushed under the skin from the insertion point to the exit site and out

- The tunneled catheter is gripped by the tool and pulled back to emerge from the insertion site.  At this point the catheter is open at both ends to the exterior environment but passes through a tunnel between the insertion point of the venous catheter and the exit point.

- The end of the tunnelled catheter that is adjacent to the venous catheter is connected to the outside end of the venous catheter. This means that there is now a continuous line from the far end of the tunnelled catheter, through the tunnel, onto the venous catheter and finally into the central vein

## Peripheral Venous Catheter (PVC)

### Common Insertion Sites for PVC

In the past  insertion sites have been defined somewhat generally e.g  Dorsum of left hand. However, this level of documentation is now deprecated and it is recommended to make a specific identification of the vein used.  Some veins traverse more than one surface site so it may be helpful for the user to be able to identify the general site and then the vein.

Scalp

- Posterior auricular

- Supratrochlear

Neck

- External jugular

Upper arm

- Brachial

- Basilic

- Cephalic

Antecubital Fossa (inner elbow region) (all gauges are suitable)

- Cephalic

- Median cephalic

- Accessory cephalic

- Basilic

- Median basilic

- Superficial median

- Median cubital

Forearm

- Cephalic (22G – 18G)

- Median antebrachial

- Accessory cephalic

Dorsum of hand

- Dorsal digital

- Dorsal metacarpal (22G)

- Dorsal venous arch

- Basilic (22G – 18G)

Umbilical region

- Umbilical (normally only used in paediatric cases)

Leg

- Subclavian vein

- Axillary vein

- Femoral vein

- Other [specify

Ankle

- Saphenous

Foot

- Great saphenous

- Lesser saphenous

- Dorsal venous arch

## Peripherally Inserted Central Venous Catheter

### Insertion site

- Cephalic vein

- Basilic vein

- Brachial vein

- Other [specify]

### Adjunctive technique

- Ultrasound

- Chest X-ray (CXR)

- Flouroscopy

- Other [specify]

## Arterial lines

Arterial lines are usually used to obtain samples of blood, typically for blood gas analysis and for invasive blood pressure measurement.

### Insertion Sites for arterial lines

- Radial artery (most common, preferred)

- Ulnar

- Brachial

- Axillary

- Dorsalis pedis

- Posterial tibial

- Femoral

### Typical Gauges for arterial lines

- 20G  adults

- 22G paediatrics

- 24G neonates and small babies

## Intra-osseous

Intraosseous infusion is particularly used with children but may also be used in adults.  There are special intraosseous needles but hypodermics and spinal needles may also sometimes be used.

### Bones used for intra-osseous lines

- Sternum

- Humerus

- Distal femur

- Proximal tibia (most common, preferred)

- Distal tibia

## Pulmonary wedge pressure (PWP)

The pulmonary capillary wedge pressure or PCWP (also called the pulmonary wedge pressure or PWP, or pulmonary artery occlusion pressure or PAOP) is the pressure measured by wedging a pulmonary catheter with an inflated balloon into a small pulmonary arterial branch.

The flow-directed balloon-tipped pulmonary artery catheter (PAC) (also known as the Swan-Ganz or right heart catheter) has been in clinical use for more than 30 years.

Typically, a multilumen catheter, 110 cm long, with extra connecting tubes for attachment to the pressure transducer. A description (example, is below).

At the tip is the PA lumen, or distal lumen. A 1.5-cc balloon is located just proximal to the tip. Approximately 4 cm proximal to the balloon is the thermistor used to measure temperature changes for calculation of CO.

Two additional lumens usually are present at 19 cm and 30 cm from the tip. Depending on the degree of right heart enlargement and the position of the catheter (ie, distance advanced into the patient), these lumina reside within the right ventricle (RV), right atrium (RA), or the superior vena cava (SVC).

Some catheters are coated with heparin to reduce thrombogenicity and have connections for temporary ventricular pacing. The former is important to remember in case the patient develops heparin-induced thrombocytopenia, because only a small amount of heparin is necessary to sustain this process.

Proper attachment of the PAC to the monitoring equipment is essential for accurate measurements.

Transmission of pressures from the body to the display system is accomplished via semirigid, noncompliant tubing filled with fluid, usually isotonic saline with a small amount of heparin. This, in turn, is connected to a fluid-filled pressure transducer.

Often, a constant infusion or "interflow" device is placed into the connecting pressure line. This device does not alter the pressure and provides a small constant infusion of fluid through the catheter to prevent backup of blood. Because the fluid is incompressible and the tubing noncompliant, this system fairly accurately transmits intracardiac pressures to the transducer, causing small amounts of movement in the transducer membrane. Deformation of this membrane generates a proportional electric current that is amplified and transmitted to the monitor.

A guide wire catheter may be useful in patients with large RVs and a lot of tricuspid regurgitation.

### PAC insertion

The PAC is inserted percutaneously into a major vein (jugular, subclavian, femoral) via an introducer sheath. The actual venous access techniques are not described here, but the following points are important. Preference considerations for cannulation of the great veins are as follows:

- Right internal jugular vein (RIJ) - Shortest and straightest path to the heart

- Left subclavian - Does not require the PAC to pass and course at an acute angle to enter the SVC (compared to the right subclavian or left internal jugular [LIJ])

- Femoral veins - These access points are distant sites, from which passing a PAC into the heart can be difficult, especially if the right-sided cardiac chambers are enlarged. Often, fluoroscopic assistance is necessary. Nevertheless, these sites are compressible and may be preferable if the risk of hemorrhage is high.

As with any catheterization procedure, sterile technique is essential. The total length of a PAC is approximately 150 cm; extra sterile towels around the head, shoulders, and chest ensure that aseptic technique is not compromised.

While the Trendelenburg position is used for venous access (internal jugular [IJ] and subclavian routes), passage of the PAC is easier when the patient subsequently is placed flat or slightly upright.

Before insertion, check the PAC for cracks and kinks. Then, check balloon function (see image below), connect all lumens to stopcocks, and flush them to eliminate air bubbles. Flick the PAC tip to check frequency response. Finally, the PAC is threaded through a sterile sleeve (be sure to check orientation) to ensure sterility of the PAC after insertion and allow some adjustment of position.

### PAC insertion

The PAC is inserted percutaneously into a major vein (jugular, subclavian, femoral) via an introducer sheath. The actual venous access techniques are not described here, but the following points are important. Preference considerations for cannulation of the great veins are as follows:

- Right internal jugular vein (RIJ) - Shortest and straightest path to the heart

- Left subclavian - Does not require the PAC to pass and course at an acute angle to enter the SVC (compared to the right subclavian or left internal jugular [LIJ])

- Femoral veins - These access points are distant sites, from which passing a PAC into the heart can be difficult, especially if the right-sided cardiac chambers are enlarged. Often, fluoroscopic assistance is necessary. Nevertheless, these sites are compressible and may be preferable if the risk of hemorrhage is high.

As with any catheterization procedure, sterile technique is essential. The total length of a PAC is approximately 150 cm; extra sterile towels around the head, shoulders, and chest ensure that aseptic technique is not compromised.

While the Trendelenburg position is used for venous access (internal jugular [IJ] and subclavian routes), passage of the PAC is easier when the patient subsequently is placed flat or slightly upright.

Before insertion, check the PAC for cracks and kinks. Then, check balloon function (see image below), connect all lumens to stopcocks, and flush them to eliminate air bubbles. Flick the PAC tip to check frequency response. Finally, the PAC is threaded through a sterile sleeve (be sure to check orientation) to ensure sterility of the PAC after insertion and allow some adjustment of position.

The packaging of the PAC causes it to have a preformed curve. This can be used to facilitate passage into the PA. The direction in which the curl is inserted into the introducer depends on which vein is cannulated. For instance, from the head of the bed using the RIJ approach, the curl should be in the direction of the patient's left shoulder (concave-cephalad). Once the PAC is in the RV, a clockwise quarter turn moves the tip anteriorly to allow easier passage into the PA.

After inserting the PAC as far as the 20-cm mark (30-cm mark if the femoral route used), the balloon is inflated with air. Inflation should be slow and controlled (1 mL/s) and should not surpass the recommended volume (usually 1.5 mL). Always inflate the balloon before advancing the PAC, and always deflate the balloon before withdrawing the PAC.

Always use continuous pressure monitoring from the distal lumen. Watch the monitor for changes in the waveform and abnormal cardiac rhythms. From the RIJ approach, the RA is entered at approximately 25 cm, the RV at approximately 30 cm, and the PA at approximately 40 cm; the PCWP can be identified at approximately 45 cm.

If an RV waveform still present approximately 20 cm after the initial RV pattern appears, the catheter may be coiling in the RV. If withdrawal is necessary, always proceed slowly to decrease the risk of knotting the catheter upon itself. If the catheter is knotted, fluoroscopy may be necessary to visualize the catheter and remove the knot. As a last resort, slowly withdraw the PAC to the point where it catches on the introducer tip. From this point, the PAC and introducer can be removed as one unit. Apply prompt pressure for a minimum of 5 minutes. If bleeding persists, suturing the site may be necessary.

Once the PCWP is obtained and the catheter sleeve secured, make sure the PCWP pattern is reproducible before removing the sterile field. Also, determine the volume of air in the balloon required to obtain a PCWP waveform. Volumes less than half the balloon maximum may indicate that the tip is too far distal. Some clinicians advocate that, after establishing that the PA diastolic pressure is equal to the PCWP pressure, further balloon inflations are unnecessary and the PA diastolic pressure should be used as the parameter to assess left ventricular (LV) filling; this relationship may not hold if the clinical situation changes.

Once the procedure is complete, obtain a chest radiograph to check the position of the PAC and to assess for central venous access complications (eg, pneumothorax).

## Routine intra-arterial pressure measurement

Common arteries to use:

Radial artery / femoral artery (90% of the time)

But also …

Brachial artery

Axillary artery

Dorsalis pedis artery

- Intra-arterial

- Intra-osseous

- Portal catheter

- Gastric tube

A description of a line and its placement will include:

- The general type of the line e.g. Peripheral IV, Central IV

- The physical components of the line e.g. cannula, catheter

- The technique of line placement e.g. ultrasound guidance

- Ancillary procedural details such as skin preparation and securing of the line in place

[Intraosseous access reference: http://reference.medscape.com/article/80431-overview]

All lines are, by definition, parenteral. Strictly speaking, parenteral is the introduction of nutrition, a medication, or other substance into the body via any route other than the mouth but this isn’t very sophisticated. For example, topical administration routes such as the eyes, nose, ears and skin are not parenteral but neither or they via the mouth. A more practical definition might be any way that involves an invasion of the body surface which includes lines but also straightforward subcutaneous and intramuscular injections.

. is defined. It therefore makes sense to organize things so that a line is established and described and then a drug administration is associated with that line. This covers:

- Peripheral IV lines

- Central IV lines

- Intra-arterial lines

- Intra-osseous lines

- Spinals and epidurals

- Portal catheter

- Gastric tube

Next we have ad hoc injections and transdermal patches that have some arbitrary anatomical site. These include:

- Subcutaneous

- Intramuscular

- Epicutaneous (administration via a needle to the surface of the skin)

- Transdermal (patches)

An epicutaneous injection is classified as a topical route but because it has a definite site, it is grouped with the subcutaneous and intramuscular which are parenteral routes. In the same way, a transdermal patch would be described as topical but also has a site that may need to be recorded.

We’re now left with a number of other options where the anatomical site is implicit viz.

- Ocular (eye)

- Auricular / Otic (ear)

- Nasal

- Buccal

- Rectal

- Vaginal

The FDA and other sources define more, and more detailed, routes e.g. Conjunctival (which is a sub-class of Ocular) and Extra-Amniotic (administration to the outside of the membrane enveloping the fetus). However, we would be in danger of totally overloading the average user if all of these were included so they are not included. Users have the option to annotate a drug administration of they wish to capture some of these more esoteric methods and should be encouraged to come back to us with feedback if they think they should be incorporated as standard options.

| Route | Class of route | Note |
| --- | --- | --- |
| Cutaneous |  |  |
| Epicutaneous | Topical | Application onto the skin). It can be used both for local effect as in allergy testing and typical local anesthesia, as well as systemic effects when the active substance diffuses through skin in a transdermal route |
| Intradermal | Topical | Into the skin itself |
| Subcutaneous | Parenteral | Under the skin |
| Nasal |  | Through the nose |
| Intravenous | Parenteral | Into a vein |
| Intraarterial |  | Into an artery |
| Intramuscular |  | Into a muscle |
| Intracardiac |  | Into the heart |
| Intraosseous |  | Into a bone |
| Intrathecal |  | Into the spinal canal (subarachnoid) |
| Intraperitoneal |  | Infusion or injection into the peritoneum |
| Intravesical |  | Infusion is into the urinary bladder |
| Intravitreal |  | Into the eye |
| Intracavernous |  | An injection into the base of the penis |
| Intravaginal |  | Into the vagina |
| Intravenous |  | Into a vein |
| Intravenous |  | Into a vein |
| Intravenous |  | Into a vein |
| Intrauterine |  |  |
| Extraamniotic |  | Between the endometrium and fetal membranes |
| Transdermal |  | Diffusion through the intact skin for systemic rather than topical distribution |
| Transmucosal |  | diffusion through a mucous membrane, sublingual, vaginal suppositories |

injection,

l administration, in the vagina

Intrauterine

[Topical]

l

Subcutaneous

- epicutaneous or topical.

- intradermal, () is used for skin testing some allergens, and also for mantoux test for Tuberculosis

- subcutaneous (under the skin), e.g. insulin. Skin popping is a slang term that includes this method of administration, and is usually used in association with recreational drugs.

- nasal administration (through the nose) can be used for topically acting substances, as well as for insufflation of e.g. decongestant nasal sprays to be taken up along the respiratory tract. Such substances are also called inhalational, e.g. inhalational anesthetics.

- intravenous (into a vein), e.g. many drugs, total parenteral nutrition

- intraarterial (into an artery), e.g. vasodilator drugs in the treatment of vasospasm and thrombolytic drugs for treatment of embolism

- intramuscular (into a muscle), e.g. many vaccines, antibiotics, and long-term psychoactive agents. Recreationally the colloquial term 'muscling' is used.[8]

- intracardiac (into the heart), e.g. adrenaline during cardiopulmonary resuscitation (no longer commonly performed)

- intraosseous infusion (into the bone marrow) is, in effect, an indirect intravenous access because the bone marrow drains directly into the venous system. This route is occasionally used for drugs and fluids in emergency medicine and pediatrics when intravenous access is difficult.

- intrathecal (into the spinal canal) is most commonly used for spinal anesthesia and chemotherapy

- intraperitoneal, (infusion or injection into the peritoneum) e.g. peritoneal dialysis

- Intravesical infusion is into the urinary bladder.

- intravitreal, through the eye

- Intracavernous injection, an injection into the base of the penis

- Intravaginal administration, in the vagina

- Intrauterine

  - Extra-amniotic administration, between the endometrium and fetal membranes

- transdermal (diffusion through the intact skin for systemic rather than topical distribution), e.g. transdermal patches such as fentanyl in pain therapy, nicotine patches for treatment of addictionand nitroglycerine for treatment of angina pectoris.

- Transmucosal (diffusion through a mucous membrane), e.g. insufflation (snorting) of cocaine, sublingual, i.e. under the tongue, sublabial, i.e. between the lips and gingiva, nitroglycerine, vaginal suppositories

A separate area should provide options to record the following:

- Nasogastric Tube (NG)

- Nasojejunal (NJ)

- Gastric drainage tube (G-tube)

Intravenous

Intra-osseous

Intra-arterial

Subcutaneous

Intra-muscular

Inhalation

Nasal

Oral

Buccal

Sub-lingual

Rectal

Vaginal

Intrauterine

Extra-amniotic

Intra-vesical

Intra-cavernous

Urethral

Ocular (eye)

Intra-vitreal

Auricular / Otic (ear)

Transdermal

Iontophoresis

Phonophoresis